FMC-5CH-POF
hardware manual

Beam Diagnostics Group (DIG)

CERN Open Hardware Licence v1.2
Through this CERN Open Hardware Licence ("CERN OHL") version 1.2, CERN wishes to provide a tool to foster collaboration and sharing among hardware designers. The CERN OHL is copyright CERN. Anyone is welcome to use the CERN OHL, in unmodified form only, for the distribution of their own Open Hardware designs. Any other right is reserved. Release of hardware designs under the CERN OHL does not constitute an endorsement of the licensor or its designs nor does it imply any involvement by CERN in the development of such designs.

Beam Diagnostics Group
Brazilian Synchrotron Light Laboratory
Tel.: +55 19 3512-5071
Email: dig@lnls.br
# Introduction
The FMC-5CH-POF is a Vita 57 FMC LPC board, with 5 Versatile Link plastic optical fiber connectors, which can be either transmitters or receivers, depending on mounting options. The communication with the carrier board is done via LVDS signals, on either direction, and the direction of each channel is sent to the carrier through a single-ended signal.
# FMC connector signals
The board is low-pin-count (LPC) FMC card, following the Vita 57 Standard. Only the 3.3V power is used in the board, with the 3.3 VAux being used for the EEPROM. The 12V power rail is not used in this board, and left floating, while Vadj is only used to adjust signal levels. Vadj can, then, be set to any value within the FMC standard with no risk or loss of functionality.

The signal I/Os are LVDS, terminated with 100 Ohms inside the FMC card. To ensure appropriate signals when sending data to the FMC connector, the LVDS transceiver supports M-LVDS. The carrier may, then, act as an usual LVDS receiver, terminated with 100 Ohms.

Besides the data channels, the FMC-5CH-POF also has 5 direction signals in the FMC connector. These signals are natively 3.3V, but the traces have diodes connect to Vadj that ensure the output signals will never be higher than Vadj+0.6V. They are also equipped with 100 Ohms resistor in series, to avoid shorts and overcurrent.

Channel	Data signal (differential)	IsReceiver signal
1	LA00_CC	LA07_P
2	LA04	LA05_P
3	LA01_CC	LA05_N
4	LA02	LA06_P
5	LA03	LA06_N

Table 1 - Data signal pinout in the FMC connector

The mandatory Vita 57 FMC signals for EEPROM, and PRSNT_M2C_L, and VREF_A_M2C are all available. VREF_A_M2C is directly connected to the Vadj rail.

# FMC board identification
The EEPROM of this board contains data following the VITA57.1 standard. There are no custom records.
# POF connectors
The board is equipped with 5 Versatile Link connectors, which can either transmitters or receivers. The direction of the communication is defined by the board assembly and cannot be changed afterwards. Any combination of transmitters and receivers is allowed.

By the Versatile Link convention, transmitters are grey in color, while receivers are black. There are no other markers for the user for the direction of each connectors.

# Power supply

The board draws current only from the FMC connector's 3.3V and 3.3Vaux rails. The maximum estimated current for the 3.3V rail is 300mA. This board does not draws any current from Vadj and 12V rails in the FMC connector.

# Data communications
The transceivers are adequate for data communication, either asynchronous or synchronous (using separate transmitters or receivers for data and clock signal). The specified data rate for the transmitters and receivers is 50Mbd.

Testing with a loopback formed by the FMC-5CH-POF two 9 meters-long POF in SP605 at ~70Mbd, using a synchronous communication protocol, the bit error rate is less than 1*10-12 with a 99% confidence. This is more than enough for most communication applications.